## LK isi
KEMENTERIAN KESEHATAN RI
Balai Pengamanan Fasilitas Kesehatan Medan
LEMBAR KERJA PENGUJIAN
LIGHT SOURCE
I. | Data Sarana Pelayanan Kesehatan |  |  |  |  | :
| 1. Nama Sarana Pelayanan Kesehatan | [5 empty] | : | RS. SANTA MARIA PEKANBARU
| 2. Alamat | [5 empty] | :
II. | Data Alat |  | :
| 1. | Merek |  | : | BELMONT
| 2. | Type / Model |  | : | CALISTA II
| 3. | Serial Number |  | : | 48569
| 4. | Tempat Kalibrasi |  | : | Ruang Poli Gigi
| 5. | Ruang |  | : | Poli Gigi
| 6. | Tanggal Kalibrasi |  | : | 11 Desember 2022
| 8 | Petugas Kalibrasi |  | : | Fauzan Aziz, AMTE
| 9 | Tanggal Laporan |  | : | 29 Desember 2022
III. | Daftar Alat Yang Digunakan |  |  | :
| No. | Nama Alat |  |  |  |  | Merk |  |  |  | Model/Type |  |  | No. Seri
| 1 | Lux Meter |  |  |  |  | Kimo |  |  |  | LX 200 |  |  | 21110978
| 2 | Electrical Safety Analyzer |  |  |  |  | Fluke |  |  |  | ESA 615 |  |  | 4427051
| 3 | Thermohygrometer |  |  |  |  | BK Precision |  |  |  | 725 |  |  | 117C21114
| 4 | Meteran |  |  |  |  | - |  |  |  | - |  |  | -
IV. | Pengukuran Kondisi Lingkungan |  |  |  | :
| No. | Parameter |  |  |  |  | Terukur
| 1 | Temperatur Ruangan |  |  |  |  | Awal |  | : | 26.5 | °C |  | Akhir | : | 26.7 | °C
| 2 | Kelembaban Ruangan |  |  |  |  | Awal |  | : | 68.9 | %RH |  | Akhir | : | 68.7 | %RH
| 3 | Tegangan Input | [7 empty] | 220.4 | Vac
V. | Pemeriksaan Konfisi Fisik dan Fungsi Alat | [6 empty] | :
| No. | Parameter |  | Batasan Pemeriksaan | [10 empty] | Hasil Pengamatan
| 1 | Badan dan permukaan |  | selungkup utuh, bersih, terpasang ketat satu dan lainnya dan tidak ada bekas tertimpa cairan ataupun gangguan lainnya. | [10 empty] | 1
| 2 | Kotak kontak alat |  | Periksa apakah ada gangguan pada kotak kontak (AC-Power). Gerak-gerakkan kotak kontak untuk memastikan keamanannya. Goyang-goyangkan kotak kontak untuk memastikan tidak ada baut atau mur yang longgar. | [10 empty] | 1
| 3 | Kabel catu utama |  | Periksa kabel, apakah terlihat ada kerusakan atau bagian isolasi yang terkelupas. | [10 empty] | 1
| 4 | Lampu dan Reflektor |  | Periksa kondisi lampu apakah terpasang dengan benar, berfungsi atau tidak. Periksa juga kaca reflektor | [10 empty] | 1
| 5 | Tombol, saklar, dan kontrol |  | Sebelum menggunakan periksa posisinya, apakah posisinya sudah benar | [10 empty] | 1
| 6 | Tombol fungsi Light Source |  | Pastikan tombol tombol tersebut berfungsi dengan sempurna | [10 empty] | 1
| 7 | Tampilan dan Indikator |  | Pastikan lampu indikator dan tampilan berfungsi seluruhnya, yakinkan bahwa bagian tampilan digital berfungsi | [10 empty] | 1
VI. | Pengukuran Keselamatan Listrik |  |  |  |  | :
| No. | Parameter | [6 empty] | terukur std |  |  |  | Ambang Batas
| 1 | Resistansi pembumian protektif | [6 empty] | - |  | Ω |  | ≤ 0,3 Ω
| 2 | Arus bocor peralatan metode langsung | [6 empty] | 12.4 |  | µA |  | ≤ 500 µA
| 3 | Resistansi isolasi | [6 empty] | OL |  | MΩ |  | > 2 MΩ
VII. | Pengukuran Kinerja |  |  | :
| 1.1. | Pengukuran Illuminance
| No | Setting Pada Alat |  | Penunjukan standart ( lux ) | [12 empty] | Toleransi
|  | (lux) |  | 1 |  |  | 2 |  |  | 3 |  | 4 |  | 5 |  | 6 | (lux)
| 1 | Max |  | 10230 |  |  | 10230 |  |  | 10230 |  | 10230 |  | 10230 |  | 10230 | 0-150.000

## Laporan Light Source
LAPORAN PENGUJIAN LIGHT SOURCE
I. | Data Sarana Pelayanan Kesehatan | [6 empty] | :
| Nama Pemilik |  |  |  | : | RS. SANTA MARIA PEKANBARU
| Merk |  |  |  | : | BELMONT
| Model/Type |  |  |  | : | CALISTA II
| No. Seri |  |  |  | : | 48569
| Tanggal Pengujian |  |  |  | : | 11 Desember 2022
| Tempat Pengujian |  |  |  | : | Ruang Poli Gigi
| Ruangan |  |  |  | : | Poli Gigi
| Petugas Pengujian |  |  |  | : | Fauzan Aziz, AMTE
II. | Hasil Pengukuran Kondisi Lingkungan | [6 empty] | :
| 1. | Suhu |  |  |  | : ( | 26.6286 | ± | 0.5 | )°C
| 2. | Kelembaban |  |  |  | : ( | 69.7153 | ± | 2.3 | )% RH
| 3. | Sumber Tegangan |  |  |  | : ( | 220.0735 | ± | 1.2 | )Vac
III. | Hasil Pemeriksaan Konfisi Fisik dan Fungsi Alat | [7 empty] | :
| 1. | Fisik Alat |  |  | : | Baik
| 2. | Fungsi Alat |  |  | : | Baik
IV. | Daftar Alat Yang Digunakan |  |  |  |  | :
| No. | Nama Alat |  |  |  |  | Merk |  |  |  | Model/type |  |  | Nomor seri
| 1 | Lux Meter |  |  |  |  | Kimo |  |  |  | LX 200 |  |  | 21110978
| 2 | Electrical Safety Analyzer |  |  |  |  | Fluke |  |  |  | ESA 615 |  |  | 4427051
| 3 | Thermohygrometer |  |  |  |  | BK Precision |  |  |  | 725 |  |  | 117C21114
| 4 | Meteran |  |  |  |  | - |  |  |  | - |  |  | -
V. | Hasil Pengukuran Keselamatan Listrik | [6 empty] | :
| No. | Parameter | [6 empty] | Terukur |  |  | Ambang Batas
yang diijinkan
| 1 | Resistansi pembumian protektif | [6 empty] | - |  | Ω | ≤ 0,3 Ω
| 2 | Arus bocor peralatan metode langsung | [6 empty] | 12.4 |  | µA | ≤ 500 µA
| 3 | Resistansi isolasi | [6 empty] | OL |  | MΩ | > 2 MΩ
| Kesimpulan : |  |  |  |  | Aman atau tidak terjadi penyimpangan fungsional
| NB : | 1. | Tanda " - " pada resistansi pembumian protektif menandakan bahwa nilai resistansi tidak diukur
karena Keselamatan listrik pada UUT ( unit under test ) bukan kelas I
|  | 2. | Tanda " OL atau OR " menandakan bahwa nilai resistansi isolasi > 2 MΩ atau ∞ ( takterhingga)
VI. | Pengukuran Kinerja | [7 empty] | :
| 1. | Parameter Illuminance |  |  |  |  | :
| Setting Pada Alat |  |  | Penunjukan |  |  |  | Ambang Batas
| ( Klux ) |  |  | Standar (KLux) |  |  |  | yang Diijinkan
| Max |  |  | 10.9572 |  |  |  | ≥ 0,7 KLux
| Ketidakpastian Pengukuran : | [6 empty] | ± | 0.22008 | KLux
VII. | Telaah Teknis dan Kesimpulan | [5 empty] | :
| 1. | Kesimpulan |  | : | Alat dinyatakan Laik Pakai
| 2. | Saran |  | : | Lakukan pengujian ulang sesuai jadwal secara berkala
VIII. | Metode Dan Peraturan Yang Digunakan | [6 empty] | :
1. PERMENKES No. 54 Tahun 2015, Tentang Pengujian dan Kalibrasi Alat Kesehatan.
2. Keputusan Direktur Jenderal Pelayanan Kesehatan No. HK.02.02/V/0412/2020,
Tentang Metode Kerja Pengujian dan Kalibrasi Alat Kesehatan MK 089-19.
Catatan :
| o | Ketidakpastian pengukuran diestimasi pada tingkat kepercayaan 95 % K=2
[11 empty] | Medan, |  | 29 Desember 2022
Mengetahui,
Ka. Instalasi Pengujian dan Kalibrasi
SRI ALFANI. ST
NIP. 19720505 199603 2 002

## olah data
PENGOLAHAN DATA
INKUBATOR BAYI
I. | Data Sarana Pelayanan Kesehatan | [6 empty] | :
| 1. Nama Sarana Pelayanan Kesehatan | [7 empty] | : | RS. SANTA MARIA PEKANBARU
| 2. Alamat | [7 empty] | : | 0
II. | Data Alat |  |  | :
| 1. | Merek |  |  |  | : | BELMONT
| 2. | Type / Model |  |  |  | : | CALISTA II
| 3. | Serial Number |  |  |  | : | 48569
| 4. | Tempat Kalibrasi |  |  |  | : | Ruang Poli Gigi
| 5. | Ruang Kalibrasi |  |  |  | : | Poli Gigi
| 6. | Tanggal Kalibrasi |  |  |  | : | 11 Desember 2022
III. | Daftar Alat Yang Digunakan |  |  |  |  | :
| No. | Nama Alat | [7 empty] | Merk |  |  |  |  | Model/Type |  |  |  | No. Seri
| 1 | Lux Meter | [7 empty] | Kimo |  |  |  |  | LX 200 |  |  |  | 21110978
| 2 | Electrical Safety Analyzer | [7 empty] | Fluke |  |  |  |  | ESA 615 |  |  |  | 4427051
| 3 | Thermohygrometer | [7 empty] | BK Precision |  |  |  |  | 725 |  |  |  | 117C21114
| 4 | Meteran | [7 empty] | - |  |  |  |  | - |  |  |  | -
IV. | Pengukuran Kondisi Lingkungan | [5 empty] | :
| No. | Parameter | [5 empty] | awal |  |  |  | akhir |  |  |  | rata2 |  |  | Koreksi Sertifikat |  | Ktps
| 1 | Temperatur Ruangan | [5 empty] | 26.5 |  | °C |  | 26.7 |  | °C |  | 26.6 |  |  | 26.6286 |  | 0.5
| 2 | Kelembaban Ruangan | [5 empty] | 68.9 |  | %RH |  | 68.7 |  | %RH |  | 68.8 |  |  | 69.7153 |  | 2.3
| 3 | Tegangan Input | [5 empty] | 220.4 | [7 empty] | Vac |  |  | 220.0735 |  | 1.2
V. | Pemeriksaan Konfisi Fisik dan Fungsi Alat | [8 empty] | :
| No | Parameter | [6 empty] | Hasil Pengamatan |  |  |  | score |  |  |  | kesimpulan
| 1 | Badan dan permukaan | [6 empty] | baik |  |  |  | 10 |  |  |  | Baik
| 2 | Kotak kontak alat | [6 empty] | baik
| 3 | Kabel catu utama | [6 empty] | baik
| 4 | Lampu dan Reflektor | [6 empty] | baik
| 5 | Tombol, saklar, dan kontrol | [6 empty] | baik
| 6 | Tombol fungsi Light Source | [6 empty] | baik
| 7 | Tampilan dan Indikator | [6 empty] | baik
VI. | Pengukuran Keselamatan Listrik | [6 empty] | :
| No. | Parameter | [7 empty] | terukur std |  |  |  |  | Ambang Batas |  |  | hasil |  |  |  |  | score |  |  | kesimpulan
| 1 | Resistansi pembumian protektif | [7 empty] | - |  |  | Ω |  | ≤ | 0.3 | Ω | memenuhi persyaratan |  |  |  |  | 40 |  |  | Aman atau tidak terjadi penyimpangan fungsional
| 2 | Arus bocor peralatan metode langsung | [7 empty] | 12.4 |  |  | µA |  | ≤ | 500 | µA | memenuhi persyaratan
| 5 | Resistansi isolasi | [7 empty] | OL |  |  | MΩ |  | > | 2 | MΩ | memenuhi persyaratan
VII. | Pengukuran Kinerja |  |  |  |  | :
| 1. | PENGUJIAN KINERJA
| 1.1 | Parameter Illuminance
|  | No | Setting Pada Alat |  |  |  | Penunjukan standart ( lux ) | [15 empty] | Mean |  | STDV
|  |  | (lux) |  |  |  | 1 |  |  | 2 |  |  | 3 |  |  | 4 |  |  | 5 |  | 6
|  | 1 | Max |  |  |  | 10230 |  |  | 10230 |  |  | 10230 |  |  | 10230 |  |  | 10230 |  | 10230 |  | 10230 |  | 0
|  | No. | Mean terkoreksi |  | Koreksi |  |  | U95 |  |  | ∑koreksi + U95 |  |  | %(∑koreksi + U95) |  |  | Toleransi |  |  | hasil |  | score
|  | 1 | 10957.2156 |  | na |  |  | 0.22008 |  |  | na |  |  | na |  |  | 0 |  | s.d | lulus |  | 100
150000
VIII. | Telaah Teknis dan Kesimpulan
| Parameter | [5 empty] | score |  |  | Hasil |  |  | Bobot | [6 empty] | fungsi alat |  | telaah teknis |  | Kesimpulan
| 1 | Parameter Illuminance |  |  |  |  | 1 |  |  | Baik |  |  | 50 | [6 empty] | Baik |  | Dalam Batas Toleransi |  | Alat dinyatakan Laik Pakai |  | Lakukan pengujian ulang sesuai jadwal secara berkala
| No. | Parameter | [6 empty] | Bobot |  |  | Total
| 1 | Pemeriksaan Fisik dan Fungsi Alat | [6 empty] | 10 |  |  | 100
| 2 | Pengukuran Keselamatan Listrik | [6 empty] | 40
| 3 | Telaah Teknis Hasil Pengukuran | [6 empty] | 50

## Serti Lux
HASIL REKALIBRASI | [9 empty] | SUMMARY
DATA LUX METER
[10 empty] | No | SN | b | a | U95
Nama Alat |  | : | Lux Meter
Merk |  | : | Kimo | [6 empty] | 1 | 21110978 | 1.0723 | -12.69 | 2.2
Type |  | : | LX 200 | [6 empty] | 2 | 21110987 | 1.1363 | -11.7588 | 2.4
Nomor Seri |  | : | 21110978 |  |  |  | 2023-03-13 |  |  | 3 | 21110988 | 1.1979 | -11.9232 | 2.9
|  |  | 21110987 |  |  |  | 2023-03-13 |  |  | 21110978 |  | 1.0723 | -12.69 | 2.2
|  |  | 21110988 |  |  |  | 2023-03-13
Ketertelusuran |  | : | 4 | PT. Kaliman
|  |  | 1 | LK-110-IDN
|  |  | 2 | LK-032-IDN
|  |  | 3 | Puslit Metrologi LIPI
|  |  | 4 | PT. Kaliman
21110978 | [7 empty] | 21110987
No | Penunjukan Standar |  | Penunjukan alat |  | Koreksi |  | uncert | No | Penunjukan Standar |  | Penunjukan alat |  | Koreksi |  | uncert
1 | 50 | lux | 47 | lux | -3 | lux | 4.2 | 1 | 50 | lux | 50 | lux | 0 | lux | 8.5
2 | 100 | lux | 98 | lux | -2 | lux | 2.8 | 2 | 100 | lux | 103 | lux | 3 | lux | 4.5
3 | 200 | lux | 200 | lux | 0 | lux | 2.4 | 3 | 200 | lux | 211 | lux | 11 | lux | 2.8
4 | 400 | lux | 409 | lux | 9 | lux | 2.2 | 4 | 400 | lux | 441 | lux | 41 | lux | 2.3
5 | 600 | lux | 624 | lux | 24 | lux | 2.2 | 5 | 600 | lux | 667 | lux | 67 | lux | 2.7
6 | 800 | lux | 843 | lux | 43 | lux | 2.2 | 6 | 800 | lux | 898 | lux | 98 | lux | 2.5
7 | 1000 | lux | 1068 | lux | 68 | lux | 2.2 | 7 | 1000 | lux | 1127 | lux | 127 | lux | 2.4
| slope b = | 1.0723 | -12.69 | = intersept : a |  |  |  |  | slope b = | 1.1363 | -11.7588 | = intersept : a
21110988
No | Penunjukan Standar |  | Penunjukan alat |  | Koreksi |  | uncert
1 | 50 | lux | 53 | lux | 3 | lux | 6.3
2 | 100 | lux | 108 | lux | 8 | lux | 3.5
3 | 200 | lux | 222 | lux | 22 | lux | 2.6
4 | 400 | lux | 468 | lux | 68 | lux | 2.3
5 | 600 | lux | 704 | lux | 104 | lux | 4.2
6 | 800 | lux | 948 | lux | 148 | lux | 3.4
7 | 1000 | lux | 1187 | lux | 187 | lux | 2.9
| slope b = | 1.1979 | -11.9232 | = intersept : a

## esa
HASIL REKALIBRASI | [7 empty] | SUMMARY
Data Esa
[8 empty] | No | SN | b | a | U
Nama Alat |  | : | Electrical Safety Analyzer |  |  |  |  | 1 | 3827025 | 0.99969 | -0.098058 | 1.2
Merk |  | : | Fluke |  |  |  |  | 2 | 4427051 | 0.99941 | -0.19592 | 1.2
Type |  | : | ESA 615 |  |  |  |  | 3 | 2937056 | 0.97829 | 4.4079 | 1.2
[8 empty] | 4 | 2086729 | 0.99991 | 0.19408 | 1.2
[8 empty] | 5 | 3827021 | 0.99528 | 0.87282 | 1.2
Nomor Seri |  | : | 3827025 |  | 2023-03-09 |  |  | 4427051 |  | 0.99941 | -0.19592 | 1.2
|  |  | 4427051 |  | 2023-03-09
|  |  | 2937056 |  | 2023-03-09
|  |  | 2086729 |  | 2023-03-09
|  |  | 3827021 |  | 2023-03-09
Ketertelusuran |  | : | 2 | LK-032-IDN
|  |  | 1 | LK-110-IDN |  | (BPFK)
|  |  | 2 | LK-032-IDN |  | (PT. Kaliman)
|  |  | 3 | Puslit Metrologi LIPI
3827025
No | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert
1 | 150 | V | 149.81 | V | -0.19 | V | 1.2
2 | 180 | V | 179.86 | V | -0.14 | V
3 | 200 | V | 199.95 | V | -0.05 | V
4 | 220 | V | 219.81 | V | -0.19 | V
5 | 230 | V | 229.78 | V | -0.22 | V
| slope b = | 0.99969 | -0.098058 | = intersept : a
| standard deviation of slope : | 0.0011897 | 0.23567 | = standard deviation of intercept
| Coefficient of determination : r2 | 1 | 0.076365 | = standard error of regression
| Fisher F-statistics | 706069.2245 | 3 | = Degree of the regression
| Sum of square due to regression | 4117.4804 | 0.017495 | = sum of square due to residual
4427051
No | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert
1 | 150 | V | 149.72 | V | -0.28 | V | 1.2
2 | 180 | V | 179.71 | V | -0.29 | V
3 | 200 | V | 199.63 | V | -0.37 | V
4 | 220 | V | 219.74 | V | -0.26 | V
5 | 230 | V | 229.64 | V | -0.36 | V
| slope b = | 0.99941 | -0.19592 | = intersept : a
| standard deviation of slope : | 0.0008261 | 0.16364 | = standard deviation of intercept
| Coefficient of determination : r2 | 1 | 0.053025 | = standard error of regression
| Fisher F-statistics | 1463596.3703 | 3 | = Degree of the regression
| Sum of square due to regression | 4115.1214 | 0.008435 | = sum of square due to residual
2937056
No | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert
1 | 150 | V | 151.17 | V | 1.17 | V | 1.2
2 | 180 | V | 180.42 | V | 0.42 | V
3 | 200 | V | 200.32 | V | 0.32 | V
4 | 220 | V | 219.11 | V | -0.89 | V
5 | 230 | V | 229.74 | V | -0.26 | V
| slope b = | 0.97829 | 4.4079 | = intersept : a
| standard deviation of slope : | 0.0060297 | 1.1944 | = standard deviation of intercept
| Coefficient of determination : r2 | 0.99989 | 0.38703 | = standard error of regression
| Fisher F-statistics | 26322.9646 | 3 | = Degree of the regression
| Sum of square due to regression | 3943.0225 | 0.44938 | = sum of square due to residual
2086729
No | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert
1 | 150 | V | 150.17 | V | 0.17 | V | 1.2
2 | 180 | V | 180.22 | V | 0.22 | V
3 | 200 | V | 200.12 | V | 0.12 | V
4 | 220 | V | 220.21 | V | 0.21 | V
5 | 230 | V | 230.16 | V | 0.16 | V
| slope b = | 0.99991 | 0.19408 | = intersept : a
| standard deviation of slope : | 0.00072434 | 0.14349 | = standard deviation of intercept
| Coefficient of determination : r2 | 1 | 0.046494 | = standard error of regression
| Fisher F-statistics | 1905599.4773 | 3 | = Degree of the regression
| Sum of square due to regression | 4119.24 | 0.006485 | = sum of square due to residual
3827021
No | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert
1 | 150 | V | 150.14 | V | 0.14 | V | 1.2
2 | 180 | V | 180.04 | V | 0.04 | V
3 | 200 | V | 200.04 | V | 0.04 | V
4 | 220 | V | 219.62 | V | -0.38 | V
5 | 230 | V | 229.9 | V | -0.1 | V
| slope b = | 0.99528 | 0.87282 | = intersept : a
| standard deviation of slope : | 0.0024128 | 0.47795 | = standard deviation of intercept
| Coefficient of determination : r2 | 0.99998 | 0.15487 | = standard error of regression
| Fisher F-statistics | 170160.6249 | 3 | = Degree of the regression
| Sum of square due to regression | 4081.2117 | 0.071953 | = sum of square due to residual

## ktps
BUDGET KETIDAKPASTIAN
INKUBATOR BAYI
| n | 6
| resolusi UUT | 0
| Komponen | Satuan | Distribusi | U | Pembagi | v | u | c | (u.c) | (u.c)² | (u.c)⁴/v
| Pembacaan berulang | lux | normal | 0 | 2.4495 | 5 | 0 | 1 | 0 | 0 | 0
| Sertifikat Standar | lux | Segi 4 | 219.1443 | 2 | 50 | 109.5722 | 1 | 109.5722 | 12006.0574 | 2882908.3068
| Drift alat standar | lux | Segi 4 | 0.021 | 1.7321 | 50 | 0.012124 | 1 | 0.012124 | 0.000147 | 4.3218e-10
| resolusi UUT | lux | Segi 4 | 0 | 1.7321 | 50 | 0 | 1 | 0 | 0 | 0
| Jumlah | [8 empty] | 12006.0576 | 2882908.3068
| Ketidakpastian baku gabungan, Uc | [8 empty] | 109.5722
| Derajat kebebasan efektif, V eff | [8 empty] | 50
| Faktor cakupan, k-student untuk Veff dan CL 95% | [8 empty] | 2.0086
| Ketidakpastian bentangan, U=k.Uc | [8 empty] | 220.0822 | lux
[10 empty] | 0.22008 | klux
2
3
4

## serti thermohygro
HASIL REKALIBRASI | [8 empty] | SUMMARY
Data Thermohygro
[9 empty] | No | SN | Suhu |  | Kelembaban |  | U Temp | U Hygro
[11 empty] | b | a | b | a
Nama Alat |  |  | : | Thermohygrometer |  |  |  |  | 1 | 117C21114 | 1.013 | -0.31715 | 0.99281 | 1.4101 | 0.5 | 2.3
Merk |  |  | : | BK Precision |  |  |  |  | 2 | 021210-02 | 1.0189 | -0.41573 | 0.94936 | 11.827 | 0.5 | 2.3
|  |  |  | Bk Precision |  |  |  |  | 3 | 34901888 | 1.0274 | -0.23216 | 1.089 | 1.3944 | 0.5 | 2.3
Nomor Seri |  |  | : | 021210-02 | Greisinger | GFTB 100 |  | 30/1/2023 | 4 | 15301 | 0.97996 | 0.56425 | 0.98892 | 13.2098 | 0.5 | 2.3
|  |  |  | 117C21114 | BK Precision | 725 |  | 2023-03-27 | 5 | 34902416 | 1.0113 | -0.066759 | 0.97465 | 8.8448 | 0.5 | 2.8
|  |  |  | 34901888 | Greisinger | GFTB 200 |  | 2022-02-02 | 6 | 117l15116 | 0.96529 | 0.78106 | 0.99608 | 6.5245 | 0.5 | 1.3
|  |  |  | 15301 | Greisinger | GFTB 100 |  | 2023-03-02 | 7 | 117l15117 | 1.0758 | -1.0797 | 0.99608 | 6.5245 | 0.5 | 1.3
|  |  |  | 10050021 | BK Precision | 720 |  | 2023-03-02 | 8 | 10050021 | 0.99731 | 0.59063 | 1.0811 | 2.3943 | 0.5 | 2.3
|  |  |  | 116C13105 | BK Precision | 720 |  | 2023-03-02 | 9 | 116C13105 | 0.99773 | -0.091265 | 1.0811 | 2.3943 | 0.5 | 2.3
|  |  |  | - | - | -
|  |  |  | 34902416 | Greisinger | GFTB 200 |  | 30/1/2023 | 117C21114 |  | 1.013 | -0.31715 | 0.99281 | 1.4101 | 0.5 | 2.3
|  |  | : | 117l15116 | Bk Precision | 720 |  | 13/2/2023
|  |  |  | 117l15117 | Bk Precision | 720 |  | 13/2/2023
Ketertelusuran |  | : | 2 | LK-032-IDN
|  |  | 1 | LK-110-IDN |  | (BPFK)
|  |  | 2 | LK-032-IDN |  | (PT. Kaliman)
|  |  | 3 | Puslit Metrologi LIPI
Suhu | [7 empty] | 117C21114 | Kelembaban
No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert |  | No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert
1 | 15.2 | °C | 15 | °C | 0.2 | °C | 0.5 |  | 1 | 29 | % | 30 | % | -1 | % | 2.3
2 | 19.9 | °C | 20 | °C | -0.1 | °C |  |  | 2 | 38.7 | % | 40 | % | -1.3 | %
3 | 25 | °C | 25 | °C | 0 | °C |  |  | 3 | 48.7 | % | 50 | % | -1.3 | %
4 | 30 | °C | 30 | °C | 0 | °C |  |  | 4 | 59 | % | 60 | % | -1 | %
5 | 34.9 | °C | 35 | °C | -0.1 | °C |  |  | 5 | 69.3 | % | 70 | % | -0.7 | %
6 | 36.9 | °C | 37 | °C | -0.1 | °C |  |  | 6 | 79.4 | % | 80 | % | -0.6 | %
7 | 39.7 | °C | 40 | °C | -0.3 | °C |  |  | 7 | 89 | % | 90 | % | -1 | %
| slope b = | 1.013 | -0.31715 | = intersept : a | [5 empty] | slope b = | 0.99281 | 1.4101 | = intersept : a
| standard deviation of slope : | 0.0045708 | 0.13721 | = standard deviation of intercept | [5 empty] | standard deviation of slope : | 0.0044532 | 0.27769 | = standard deviation of intercept
| Coefficient of determination : r2 | 0.9999 | 0.10238 | = standard error of regression | [5 empty] | Coefficient of determination : r2 | 0.9999 | 0.23734 | = standard error of regression
| Fisher F-statistics | 49116.2432 | 5 | = Degree of the regression | [5 empty] | Fisher F-statistics | 49702.5 | 5 | = Degree of the regression
| Sum of square due to regression | 514.8047 | 0.052407 | = sum of square due to residual | [5 empty] | Sum of square due to regression | 2799.7184 | 0.28165 | = sum of square due to residual
Suhu | [7 empty] | 021210-02 | Kelembaban
No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert |  | No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert
1 | 15.2 | °C | 15 | °C | 0.2 | °C | 0.5 |  | 1 | 18.8 | % | 30 | % | -11.2 | % | 2.3
2 | 20 | °C | 20 | °C | 0 | °C |  |  | 2 | 29.7 | % | 40 | % | -10.3 | %
3 | 24.8 | °C | 25 | °C | -0.2 | °C |  |  | 3 | 40.4 | % | 50 | % | -9.6 | %
4 | 30 | °C | 30 | °C | 0 | °C |  |  | 4 | 51 | % | 60 | % | -9 | %
5 | 34.7 | °C | 35 | °C | -0.3 | °C |  |  | 5 | 61.5 | % | 70 | % | -8.5 | %
6 | 36.7 | °C | 37 | °C | -0.3 | °C |  |  | 6 | 71.8 | % | 80 | % | -8.2 | %
7 | 39.7 | °C | 40 | °C | -0.3 | °C |  |  | 7 | 82 | % | 90 | % | -8 | %
| slope b = | 1.0189 | -0.41573 | = intersept : a | [5 empty] | slope b = | 0.94936 | 11.827 | = intersept : a
| standard deviation of slope : | 0.0047653 | 0.14265 | = standard deviation of intercept | [5 empty] | standard deviation of slope : | 0.0047466 | 0.26079 | = standard deviation of intercept
| Coefficient of determination : r2 | 0.99989 | 0.10611 | = standard error of regression | [5 empty] | Coefficient of determination : r2 | 0.99988 | 0.26455 | = standard error of regression
| Fisher F-statistics | 45721.3758 | 5 | = Degree of the regression | [5 empty] | Fisher F-statistics | 40002.7645 | 5 | = Degree of the regression
| Sum of square due to regression | 514.8008 | 0.056298 | = sum of square due to residual | [5 empty] | Sum of square due to regression | 2799.6501 | 0.34993 | = sum of square due to residual
Suhu | [7 empty] | 34901888 | Kelembaban
No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert |  | No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert
1 | 14.6 | °C | 15 | °C | -0.4 | °C | 0.5 |  | 1 | 26.1 | % | 30 | % | -3.9 | % | 2.3
2 | 19.7 | °C | 20 | °C | -0.3 | °C |  |  | 2 | 35.5 | % | 40 | % | -4.5 | %
3 | 24.8 | °C | 25 | °C | -0.2 | °C |  |  | 3 | 44.7 | % | 50 | % | -5.3 | %
4 | 29.6 | °C | 30 | °C | -0.4 | °C |  |  | 4 | 53.9 | % | 60 | % | -6.1 | %
5 | 34.4 | °C | 35 | °C | -0.6 | °C |  |  | 5 | 63.1 | % | 70 | % | -6.9 | %
6 | 36.2 | °C | 37 | °C | -0.8 | °C |  |  | 6 | 72.2 | % | 80 | % | -7.8 | %
7 | 38.9 | °C | 40 | °C | -1.1 | °C |  |  | 7 | 81.2 | % | 90 | % | -8.8 | %
| slope b = | 1.0274 | -0.23216 | = intersept : a | [5 empty] | slope b = | 1.089 | 1.3944 | = intersept : a
| standard deviation of slope : | 0.009752 | 0.28787 | = standard deviation of intercept | [5 empty] | standard deviation of slope : | 0.0028288 | 0.16085 | = standard deviation of intercept
| Coefficient of determination : r2 | 0.99955 | 0.21533 | = standard error of regression | [5 empty] | Coefficient of determination : r2 | 0.99997 | 0.13745 | = standard error of regression
| Fisher F-statistics | 11098.5103 | 5 | = Degree of the regression | [5 empty] | Fisher F-statistics | 148207.1973 | 5 | = Degree of the regression
| Sum of square due to regression | 514.6253 | 0.23184 | = sum of square due to residual | [5 empty] | Sum of square due to regression | 2799.9055 | 0.094459 | = sum of square due to residual
Suhu | [7 empty] | 15301 | Kelembaban
No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert |  | No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert
1 | 14.8 | °C | 15 | °C | -0.2 | °C | 0.4 |  | 1 | 17.3 | % | 30 | % | -12.7 | % | 2.5
2 | 19.8 | °C | 20 | °C | -0.2 | °C |  |  | 2 | 27.1 | % | 40 | % | -12.9 | %
3 | 24.9 | °C | 25 | °C | -0.1 | °C |  |  | 3 | 37 | % | 50 | % | -13 | %
4 | 30 | °C | 30 | °C | 0 | °C |  |  | 4 | 47.1 | % | 60 | % | -12.9 | %
5 | 35.1 | °C | 35 | °C | 0.1 | °C |  |  | 5 | 57.2 | % | 70 | % | -12.8 | %
6 | 37.2 | °C | 37 | °C | 0.2 | °C |  |  | 6 | 67.5 | % | 80 | % | -12.5 | %
7 | 40.3 | °C | 40 | °C | 0.3 | °C |  |  | 7 | 78 | % | 90 | % | -12 | %
| slope b = | 0.97996 | 0.56425 | = intersept : a | [5 empty] | slope b = | 0.98892 | 13.2098 | = intersept : a
| standard deviation of slope : | 0.0022276 | 0.067205 | = standard deviation of intercept | [5 empty] | standard deviation of slope : | 0.0049966 | 0.2571 | = standard deviation of intercept
| Coefficient of determination : r2 | 0.99997 | 0.051579 | = standard error of regression | [5 empty] | Coefficient of determination : r2 | 0.99987 | 0.26734 | = standard error of regression
| Fisher F-statistics | 193520.2289 | 5 | = Degree of the regression | [5 empty] | Fisher F-statistics | 39171.8524 | 5 | = Degree of the regression
| Sum of square due to regression | 514.8438 | 0.013302 | = sum of square due to residual | [5 empty] | Sum of square due to regression | 2799.6426 | 0.35735 | = sum of square due to residual
Suhu | [7 empty] | 34902416 | Kelembaban
No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert |  | No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert
1 | 14.9 | °C | 15 | °C | -0.1 | °C | 0.5 |  | 1 | 20.8 | % | 30 | % | -9.2 | % | 2.8
2 | 19.8 | °C | 20 | °C | -0.2 | °C |  |  | 2 | 32 | % | 40 | % | -8 | %
3 | 24.7 | °C | 25 | °C | -0.3 | °C |  |  | 3 | 42.8 | % | 50 | % | -7.2 | %
4 | 30 | °C | 30 | °C | 0 | °C |  |  | 4 | 53.2 | % | 60 | % | -6.8 | %
5 | 34.6 | °C | 35 | °C | -0.4 | °C |  |  | 5 | 63.3 | % | 70 | % | -6.7 | %
6 | 36.6 | °C | 37 | °C | -0.4 | °C |  |  | 6 | 73 | % | 80 | % | -7 | %
7 | 39.6 | °C | 40 | °C | -0.4 | °C |  |  | 7 | 82.3 | % | 90 | % | -7.7 | %
| slope b = | 1.0113 | -0.066759 | = intersept : a | [5 empty] | slope b = | 0.97465 | 8.8448 | = intersept : a
| standard deviation of slope : | 0.0060657 | 0.18094 | = standard deviation of intercept | [5 empty] | standard deviation of slope : | 0.013682 | 0.771 | = standard deviation of intercept
| Coefficient of determination : r2 | 0.99982 | 0.13608 | = standard error of regression | [5 empty] | Coefficient of determination : r2 | 0.99902 | 0.74245 | = standard error of regression
| Fisher F-statistics | 27798.3906 | 5 | = Degree of the regression | [5 empty] | Fisher F-statistics | 5074.4825 | 5 | = Degree of the regression
| Sum of square due to regression | 514.7646 | 0.092589 | = sum of square due to residual | [5 empty] | Sum of square due to regression | 2797.2438 | 2.7562 | = sum of square due to residual
Suhu | [7 empty] | 117l15116 | Kelembaban
No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert |  | No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert
1 | 15.2 | °C | 15 | °C | 0.2 | °C | 0.5 |  | 1 | 23.6 | % | 30 | % | -6.4 | % | 1.3
2 | 19.7 | °C | 20 | °C | -0.3 | °C |  |  | 2 | 33.6 | % | 40 | % | -6.4 | %
3 | 24.8 | °C | 25 | °C | -0.2 | °C |  |  | 3 | 43.6 | % | 50 | % | -6.4 | %
4 | 30 | °C | 30 | °C | 0 | °C |  |  | 4 | 53.7 | % | 60 | % | -6.3 | %
5 | 35.4 | °C | 35 | °C | 0.4 | °C |  |  | 5 | 63.7 | % | 70 | % | -6.3 | %
6 | 37.6 | °C | 37 | °C | 0.6 | °C |  |  | 6 | 73.8 | % | 80 | % | -6.2 | %
7 | 40.9 | °C | 40 | °C | 0.9 | °C |  |  | 7 | 83.8 | % | 90 | % | -6.2 | %
| slope b = | 0.96529 | 0.78106 | = intersept : a | [5 empty] | slope b = | 0.99608 | 6.5245 | = intersept : a
| standard deviation of slope : | 0.013036 | 0.39645 | = standard deviation of intercept | [5 empty] | standard deviation of slope : | 0.00061376 | 0.035179 | = standard deviation of intercept
| Coefficient of determination : r2 | 0.99909 | 0.3063 | = standard error of regression | [5 empty] | Coefficient of determination : r2 | 1 | 0.032605 | = standard error of regression
| Fisher F-statistics | 5482.7178 | 5 | = Degree of the regression | [5 empty] | Fisher F-statistics | 2633907 | 5 | = Degree of the regression
| Sum of square due to regression | 514.388 | 0.4691 | = sum of square due to residual | [5 empty] | Sum of square due to regression | 2799.9947 | 0.0053153 | = sum of square due to residual
Suhu | [7 empty] | 117l15117 | Kelembaban
No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert |  | No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert
1 | 14.8 | °C | 15 | °C | -0.2 | °C | 0.5 |  | 1 | 23.6 | % | 30 | % | -6.4 | % | 1.3
2 | 19.5 | °C | 20 | °C | -0.5 | °C |  |  | 2 | 33.6 | % | 40 | % | -6.4 | %
3 | 24.1 | °C | 25 | °C | -0.9 | °C |  |  | 3 | 43.6 | % | 50 | % | -6.4 | %
4 | 30 | °C | 30 | °C | 0 | °C |  |  | 4 | 53.7 | % | 60 | % | -6.3 | %
5 | 33.3 | °C | 35 | °C | -1.7 | °C |  |  | 5 | 63.7 | % | 70 | % | -6.3 | %
6 | 35.2 | °C | 37 | °C | -1.8 | °C |  |  | 6 | 73.8 | % | 80 | % | -6.2 | %
7 | 37.9 | °C | 40 | °C | -2.1 | °C |  |  | 7 | 83.8 | % | 90 | % | -6.2 | %
| slope b = | 1.0758 | -1.0797 | = intersept : a | [5 empty] | slope b = | 0.99608 | 6.5245 | = intersept : a
| standard deviation of slope : | 0.027602 | 0.79892 | = standard deviation of intercept | [5 empty] | standard deviation of slope : | 0.00061376 | 0.035179 | = standard deviation of intercept
| Coefficient of determination : r2 | 0.99672 | 0.58124 | = standard error of regression | [5 empty] | Coefficient of determination : r2 | 1 | 0.032605 | = standard error of regression
| Fisher F-statistics | 1518.9628 | 5 | = Degree of the regression | [5 empty] | Fisher F-statistics | 2633907 | 5 | = Degree of the regression
| Sum of square due to regression | 513.1679 | 1.6892 | = sum of square due to residual | [5 empty] | Sum of square due to regression | 2799.9947 | 0.0053153 | = sum of square due to residual
Suhu | [7 empty] | 10050021 | Kelembaban
No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert |  | No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert
1 | 14.4 | °C | 15 | °C | -0.6 | °C | 0.5 |  | 1 | 25.6 | % | 30 | % | -4.4 | % | 2.3
2 | 19.5 | °C | 20 | °C | -0.5 | °C |  |  | 2 | 34.8 | % | 40 | % | -5.2 | %
3 | 24.5 | °C | 25 | °C | -0.5 | °C |  |  | 3 | 44 | % | 50 | % | -6 | %
4 | 29.5 | °C | 30 | °C | -0.5 | °C |  |  | 4 | 53.2 | % | 60 | % | -6.8 | %
5 | 34.5 | °C | 35 | °C | -0.5 | °C |  |  | 5 | 62.5 | % | 70 | % | -7.5 | %
6 | 36.5 | °C | 37 | °C | -0.5 | °C |  |  | 6 | 71.8 | % | 80 | % | -8.2 | %
7 | 39.5 | °C | 40 | °C | -0.5 | °C |  |  | 7 | 81.1 | % | 90 | % | -8.9 | %
| slope b = | 0.99731 | 0.59063 | = intersept : a | [5 empty] | slope b = | 1.0811 | 2.3943 | = intersept : a
| standard deviation of slope : | 0.0013641 | 0.040402 | = standard deviation of intercept | [5 empty] | standard deviation of slope : | 0.0013461 | 0.075927 | = standard deviation of intercept
| Coefficient of determination : r2 | 0.99999 | 0.031035 | = standard error of regression | [5 empty] | Coefficient of determination : r2 | 0.99999 | 0.065886 | = standard error of regression
| Fisher F-statistics | 534540.6341 | 5 | = Degree of the regression | [5 empty] | Fisher F-statistics | 645009.6154 | 5 | = Degree of the regression
| Sum of square due to regression | 514.8523 | 0.0048158 | = sum of square due to residual | [5 empty] | Sum of square due to regression | 2799.9783 | 0.021705 | = sum of square due to residual
Suhu | [7 empty] | 116C13105 | Kelembaban
No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert |  | No. | Penunjukan Standar |  | Penunjukan Alat |  | Koreksi |  | uncert
1 | 15.1 | °C | 15 | °C | 0.1 | °C | 0.5 |  | 1 | 25.6 | % | 30 | % | -4.4 | % | 2.3
2 | 20.1 | °C | 20 | °C | 0.1 | °C |  |  | 2 | 34.8 | % | 40 | % | -5.2 | %
3 | 25.2 | °C | 25 | °C | 0.2 | °C |  |  | 3 | 44 | % | 50 | % | -6 | %
4 | 30.2 | °C | 30 | °C | 0.2 | °C |  |  | 4 | 53.2 | % | 60 | % | -6.8 | %
5 | 35.2 | °C | 35 | °C | 0.2 | °C |  |  | 5 | 62.5 | % | 70 | % | -7.5 | %
6 | 37.2 | °C | 37 | °C | 0.2 | °C |  |  | 6 | 71.8 | % | 80 | % | -8.2 | %
7 | 40.1 | °C | 40 | °C | 0.1 | °C |  |  | 7 | 81.1 | % | 90 | % | -8.9 | %
| slope b = | 0.99773 | -0.091265 | = intersept : a | [5 empty] | slope b = | 1.0811 | 2.3943 | = intersept : a
| standard deviation of slope : | 0.0023661 | 0.071598 | = standard deviation of intercept | [5 empty] | standard deviation of slope : | 0.0013461 | 0.075927 | = standard deviation of intercept
| Coefficient of determination : r2 | 0.99997 | 0.053808 | = standard error of regression | [5 empty] | Coefficient of determination : r2 | 0.99999 | 0.065886 | = standard error of regression
| Fisher F-statistics | 177818.6607 | 5 | = Degree of the regression | [5 empty] | Fisher F-statistics | 645009.6154 | 5 | = Degree of the regression
| Sum of square due to regression | 514.8427 | 0.014477 | = sum of square due to residual | [5 empty] | Sum of square due to regression | 2799.9783 | 0.021705 | = sum of square due to residual

## Drift (hidden)
DRIFT = Koreksi Tahun ini - Koreksi Tahun Sebelumnya
DRIFT | [16 empty] | Ket : | Belum ada sertifikatnya
No | SN | Titik Ukur | Tahun Sebelumnya 2022 | [5 empty] | Tahun Ini 2023 | [5 empty] | 2022 | 2023
|  |  | T1 | T2 | T3 | T4 | T5 | TC | T1 | T2 | T3 | T4 | T5 | TC | Temp. Heater
1 | 43580040 | 32 | -0.12 | -0.04 | -0.17 | -0.2 | -0.31 | -0.5 | -0.12 | -0.04 | -0.17 | -0.2 | -0.31 | -0.5
|  | 36 | -0.09 | -0.04 | -0.13 | -0.23 | -0.38 | -0.5 | -0.09 | -0.04 | -0.13 | -0.23 | -0.38 | -0.5 | -0.06 | -0.06
2 | 36440019 | 32 | -0.01 | -0.22 | -0.13 | -0.48 | -0.16 | 0 | -0.01 | -0.22 | -0.13 | -0.48 | -0.16 | 0
|  | 36 | 0.01 | -0.35 | -0.14 | -0.71 | -0.14 | 0 | 0.01 | -0.35 | -0.14 | -0.71 | -0.14 | 0 | -0.02 | -0.02
3 | 36440008 | 32 | 0.06 | -0.08 | 0.28 | 0.05 | 0.1 | -0.06 | -0.23 | -0.49 | -0.54 | -0.41 | -0.47 | -1.71
|  | 36 | -0.1 | -0.28 | 0.31 | -0.13 | 0 | -0.15 | -0.04 | -0.07 | -0.2 | -0.01 | -0.01 | -1.5 | 0 | -0.08
4 | 38350016 | 32 | -0.16 | -0.16 | -0.19 | -0.03 | 0.07 | 0 | -0.16 | -0.16 | -0.19 | -0.03 | 0.07 | 0
|  | 36 | -0.07 | -0.08 | -0.11 | -0.06 | 0.05 | 0 | -0.07 | -0.08 | -0.11 | -0.06 | 0.05 | 0 | 0.09 | 0.09
5 | 43160038 | 32 | 0.11 | 0.13 | -0.03 | -0.06 | -0.03 | 0 | 0.11 | 0.13 | -0.03 | -0.06 | -0.03 | 0
|  | 36 | 0.08 | 0.15 | -0.09 | -0.04 | -0.03 | 0 | 0.08 | 0.15 | -0.09 | -0.04 | -0.03 | 0 | 0 | 0
6 | 43580041 | 32 | -0.7 | -0.86 | -0.78 | -0.17 | -1.02 | -0.39 | -0.7 | -0.86 | -0.78 | -0.17 | -1.02 | -0.39
|  | 36 | -0.71 | -1.1 | -1.03 | -0.19 | -1.32 | -0.34 | -0.71 | -1.1 | -1.03 | -0.19 | -1.32 | -0.34 | 0 | 0
7 | 43580047 | 32 | -0.49 | -0.58 | -0.22 | -0.88 | -0.48 | -0.51 | -0.09 | -0.17 | -0.04 | -0.2 | 0 | -0.32
|  | 36 | -0.76 | -0.87 | -0.31 | -1.35 | -0.94 | -0.95 | -0.07 | -0.15 | 0 | -0.19 | 0.03 | -0.42 | 0.08 | -0.14
8 | 38350009 | 32 | -0.19 | -0.12 | 0.1 | -0.17 | -0.13 | -0.3 | -0.24 | -0.3 | -0.09 | -0.19 | -0.21 | -0.7
|  | 36 | -0.13 | -0.13 | 0.3 | -0.18 | -0.04 | -0.3 | -0.11 | -0.11 | 0.02 | -0.04 | 0.08 | -0.4 | -0.04 | 0.12
9 | 43580039 | 32 | -0.07 | -0.12 | -0.16 | -0.22 | -0.29 | -0.3 | -0.07 | -0.12 | -0.16 | -0.22 | -0.29 | -0.3
|  | 36 | -0.07 | -0.07 | -0.18 | -0.24 | -0.24 | -0.3 | -0.07 | -0.07 | -0.18 | -0.24 | -0.24 | -0.3 | 0 | 0
10 | 36440025 | 32
36
11 | 38350014 | 32
36
No | SN | 32 | [5 empty] | 36 | [5 empty] | Temp Heater
|  | T1 | T2 | T3 | T4 | T5 | TC | T1 | T2 | T3 | T4 | T5 | TC
1 | 43580040 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
2 | 36440019 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
3 | 36440008 | -0.29 | -0.41 | -0.82 | -0.46 | -0.57 | -1.65 | 0.06 | 0.21 | -0.51 | 0.12 | -0.01 | -1.35 | -0.08
4 | 38350016 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
5 | 43160038 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
6 | 43580041 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
7 | 43580047 | 0.4 | 0.41 | 0.18 | 0.68 | 0.48 | 0.19 | 0.69 | 0.72 | 0.31 | 1.16 | 0.97 | 0.53 | -0.22
8 | 38350009 | -0.05 | -0.18 | -0.19 | -0.02 | -0.08 | -0.1 | 0.02 | 0.02 | -0.28 | 0.14 | 0.12 | -0.1 | 0.16
9 | 43580039 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0
36440025
38350014
21110978 |  | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A | #N/A
| Jika baru pertama kali direkal maka drift |  |  |  | : | 0